Plate reader

DESCRIPTION:

Plate reader spectrophotometer that has 340–850 nm wavelength range, enabling a wide variety of applications from enzyme kinetic studies to Lowry assays. The instrument provides fast and accurate measurements enabling complete 96-well plate reading in less than seven seconds. The instrument is also equipped with linear shaking as standard. A model with an incubator enables incubation temperatures up to 50°C and the capacity to read 384-well plates. software supports optimal use of the instrument with a visual workflow, effortless data analysis and exporting capabilities.

TECHNICAL SPECIFICATIONS:

Description			Plate reader

Warranty			Minimum 1 Year

Detector Type			Photodetector

Light Source			Quartz-halogen lamp

Reading Options		96 Well Plate; 384 Well Plate

Read-out Range		0 to 6 Abs

Software Type			Software with option to upgrade

Plate Type			96-well plates

Wavelength Selection		Filters

Wavelength Range	340 to 850nm (Flexible to upgrade), able to Scan and determine lambda max.

Filters	8-position filter wheel; filters installed: 405nm, 450nm, 540 and   620nm supply with two more filters, 690nm and 790nm (Flexible to installed new filters)

Linearity			0–3 Abs, ± 2% at 405nm, 96-well plate

384-well plate: 0 to 2.5 Abs, ±2% at 405 nm

Compatibility			Robot

Resolution			0.001 Abs

Accuracy			± 1% (0–3 Abs) or ± 0.003 Abs

Precision			CV ≤ 0.2% (0.3–3 Abs) at 405nm

Incubator Temperature	Up to 50°C

Shaking			Linear (able to shake plate during processing)

Measurement Speed	12 seconds, 96-well plate, fast mode; 13 seconds, 96-well plate, normal mode

Interface			User: On-board or PC control

PC should be provided along with the instrument for software installation.

Note: The instrument should be of country of origin with minimum of 1 year warranty. Only authorized distributor of the particular instrument will be entertained